APPENDIX     B

| Instructions: Please provide all available information requested below. ### Return to: Section: | Underwriters Laboratories Inc. NO.98 HENGSHAN RD SUZHOU NEW DISTRICT UL-CCIC Company Limited, Suzhou Branch Tel: 0512-68086400 Fax: 0512-68084099 Total Page: 18 |
| --- | --- |

| ## Wire & Cable Test Equipment List |  |
| --- | --- |
| Applicant Name & Address: | Contact Person: |
|  | Tel. No.: |
|  | Fax No.: |
|  | E-mail: |
| Manufacturer Name & Address: | Contact Person: |
|  | Tel. No.: |
|  | Fax No.: |
|  | E-mail: |
| # Type of Product(s) Produced at Manufacturing Location [ ] Flexible Power Cord [ ] Appliance Wiring Material [ ] Building Wire (18 AWG - 2000 kcmil) [ ] Communication Cable [ ] Wiring Harness [ ] Cord Sets with Molded Plug & Connectors [ ] Other ___________________________________ | Types of Manufacturing Processes Employed at Manufacturing Location: [ ] Wire Extrusion. Indicate Number of extrusion lines ____ [ ] Flat Cable Lamination [ ] Plug Molding [ ] Wire Harness Assembly [ ] Other: __________________ |
| - Does the Applicant currently have UL or CSA certification? [ ] Yes [ ] No If so, list file number and product category of each certification: | Does the Manufacturing location currently have UL or CSA certification? [ ] Yes [ ] No If so, list file number and product category of each certification: |
| - For UL Engineer’s or Field Representative’s Use: Inspection Conducted by: File No: E Inspection Date: Project No.: |  |

| Temperature Controlled Test Room, Reference: UL 1581, Sect. 440.1 |  |  |  |
| --- | --- | --- | --- |
| Criteria | Compliance with Requirements? |  | Comments |
|  | Yes | No |  |
| Test room available |  |  |  |
| Temperature Range: 24 ± 8ºC |  |  |  |
| Heating provided |  |  |  |
| Cooling provided |  |  |  |
| Thermometer available to measure room temperature |  |  |  |

| ### Optical Measuring Device Reference: UL 1581, Sect. 250.9, 240.7, 240.14, 250.5, 250.10, 260.8, 280.2 |  |  |  |
| --- | --- | --- | --- |
| Criteria | Compliance with Requirements? |  | Comments |
|  | Yes | No |  |
| Optical device available |  |  | Manufacturer: Model: Type of optical device: [ ] profile projector [ ] microscope [ ] digital imaging |
| Calibrated within the last year |  |  | Last calibration date: Next calibration date: |
| Minimum increment equal to or less than 0.001 in. (0.01 mm) |  |  | Record minimum increment: |

| Machinist Micrometer Calipers, Reference: UL 1581, Sect. 200.1, 240.1A, 250.2 |  |  |  |
| --- | --- | --- | --- |
| Criteria | Compliance with Requirements? |  | Comments |
|  | Yes | No |  |
| Machinist Micrometer Calipers available |  |  | Manufacturer: Model: |
| Calibrated within the last year |  |  | Last calibration date: Next calibration date: |
| Minimum increment equal to or less than 0.001 in. (0.01 mm) |  |  | Record minimum increment: |
| Anvil and spindle surfaces flat and clean |  |  |  |
| Micrometer registers exactly 0.000 in. when fully closed |  |  |  |

| Dead Weight Dial Micrometer Exerting 85 grf., Reference: UL 1581, Sect. 280.3, 440.24 |  |  |  |
| --- | --- | --- | --- |
| Criteria | Compliance with Requirements? |  | Comments |
|  | Yes | No | [ ] Optical device used instead. |
| Dead Weight Dial Micrometer available |  |  | Manufacturer: Model: |
| Calibrated within the last year |  |  | Last calibration date: Next calibration date: |
| Minimum increment equal to or less than 0.001 in. (0.01 mm) |  |  | Record minimum increment: |
| Presser foot diameter: 0.250 ± 0.010 in. (6.4 ± 0.2 mm) |  |  | Record actual presser foot diameter: |
| Total force exerted on specimen: 85 ± 3 grf. |  |  |  |

| Pin Gage Dial Micrometer Exerting 25 grf. Reference: UL 1581, Sect. 240.10, 250.6, 260.3 |  |  |  |
| --- | --- | --- | --- |
| Criteria | Compliance with Requirements? |  | Comments |
|  | Yes | No | [ ] Optical device used instead. |
| Dead Weight Dial Micrometer available |  |  | Manufacturer: Model: |
| Calibrated within the last year |  |  | Last calibration date: Next calibration date: |
| Minimum increment equal to or less than 0.001 in. (0.01 mm) |  |  | Record minimum increment: |
| Presser foot dimensions: 0.312 x 0.043 in. (7.92 x 1.09 mm) |  |  | Record actual presser foot dimensions: |
| Total force exerted on specimen: 25 ± 2 grf. |  |  |  |
| Pin diameter equal to or less than 0.043 in. (1.09 mm) (Note smaller diameter pins may be used to determine average and minimum thickness of wire with core diameters less than 0.043 in.) |  |  | Record actual pin diameters: |
| Pin length: 0.437 in. (11.10 mm) |  |  | Record actual pin length: |

| Dead Weight Dial Micrometer Exerting 10 grf. Reference: UL 1581, Sect. 250.2, 240.1B, 280.1 |  |  |  |
| --- | --- | --- | --- |
| Criteria | Compliance with Requirements? |  | Comments |
|  | Yes | No | [ ] Optical device used instead. |
| Dead Weight Dial Micrometer available |  |  | Manufacturer: Model: |
| Calibrated within the last year |  |  | Last calibration date: Next calibration date: |
| Minimum increment equal to or less than 0.001 in. (0.01 mm) |  |  | Record minimum increment: |
| Anvil and presser foot dimensions 0.375 x 0.078 in. (9.52 x 1.98 mm) |  |  | Record actual dimensions: Anvil – Presser foot – |
| Total force exerted on specimen: 10 ± 2 grf. |  |  |  |

| Precision Balance, Reference: UL 1581, Sect. 470.4-470.10 |  |  |  |
| --- | --- | --- | --- |
| Criteria | Compliance with Requirements? |  | Comments |
|  | Yes | No |  |
| Precision balance available |  |  | Manufacturer: Model: |
| Calibrated within the last year |  |  | Last calibration date: Next calibration date: |
| Minimum increment equal to or less than 0.1 g |  |  | Record minimum increment: |
| Set-up to allow weight of specimen to be measured in water. |  |  |  |
| Required Materials available: alcohol, distilled water, fine wire (not larger in diameter than 0.0050 in. or 0.127 mm / 36 AWG) |  |  |  |

| Tensile Tester, Reference: UL 1581, Sect. 470.1-480.9 |  |  |  |
| --- | --- | --- | --- |
| Criteria | Compliance with Requirements? |  | Comments |
|  | Yes | No |  |
| Tensile tester available |  |  | Manufacturer: Model: |
| Calibrated within the last year |  |  | Last calibration date: Next calibration date: |
| Must be located in temperature controlled test room. |  |  |  |
| Minimum increment load scale no greater than 2% value read. |  |  | Minimum increment of scale: |
| Rate of grip separation: Most materials: 20 ± 1 in./min (500 ± 25 mm/min.) |  |  |  |
| PP, SRPVC and HDPE: 2 ± 0.2 in./min. (50 ± 5 mm/min.) |  |  |  |
| Ruler with minimum increment 0.10 in. |  |  |  |
| Specimen marker: 1 in. metal blades. |  |  |  |
| Dies for cutting specimens: ASTM C: 0.250 in. |  |  |  |
| ASTM D: 0.125 in. |  |  |  |

| Dielectric Tester, Reference: UL 1581, Sect. 700-830 |  |  |  |
| --- | --- | --- | --- |
| Criteria | Compliance with Requirements? |  | Comments |
|  | Yes | No |  |
| Dielectric tester available |  |  | Manufacturer: Model: |
| Calibrated within the last year |  |  | Last calibration date: Next calibration date: |
| Range of test voltage suitable for wire types to be tested |  |  | Range: 0 - kv; minimum increment: kv Range: 0 - kv; minimum increment: kv Range: 0 - kv; minimum increment: kv |
| Test voltage 48-62 Hz, variable from 0 to at least the specified rms test potential at a maximum rate of 500 V/s). |  |  | Voltage frequency: Hz |
| Voltmeter present |  |  |  |
| Visual or audible means of indicating test failure. |  |  |  |

| Insulation Resistance Tester or Megohm Bridge, Reference: UL 1581, Sect. 920 |  |  |  |
| --- | --- | --- | --- |
| Criteria | Compliance with Requirements? |  | Comments |
|  | Yes | No |  |
| Insulation resistance tester available |  |  | Manufacturer: Model: |
| Calibrated within the last year |  |  | Last calibration date: Next calibration date: |
| Open-circuit potential: 100 - 500 V |  |  | Voltage used: |
| Resolution: 10% or less of the value indicated on meter |  |  | Minimum increment of scale: Megohms For the Ranges: Megohms respectively |
| Range applicable for requirements of wire being tested. |  |  |  |

| Water Tank for Insulation Resistance & Dielectric Tests, Reference: UL 1581, Sect. 920 |  |  |  |
| --- | --- | --- | --- |
| Criteria | Compliance with Requirements? |  | Comments |
|  | Yes | No |  |
| Water tank available |  |  |  |
| Water temperature maintained at 10.0 - 26.7ºC |  |  |  |
| Water tank equipped with copper plate electrode |  |  |  |
| Thermometer available to record water temperature |  |  | Manufacturer: Model: |
| Calibrated within the last year |  |  | Last calibration date: Next calibration date: |

| Tightness of Insulation Tester, Reference: UL 62, Sect. 17.1 |  |  |  |
| --- | --- | --- | --- |
| Criteria | Compliance with Requirements? |  | Comments [ ] Manufacturer will not produce parallel flexible cord or decorative lighting string wire or cord. |
|  | Yes | No |  |
| Tightness of insulation tester available |  |  | Manufacturer: Model: |
| Calibrated within the last year |  |  | Last calibration date: Next calibration date: |
| Clamps or other means available to securely hold cable sample under test conditions |  |  |  |
| Weight values: 4 lbf. (1.81 kg) 8 lbf. (3.63 kg) |  |  |  |
| Weight value permanently marked on weights |  |  |  |

| Mechanical Strength Tester, Reference: UL 62, Sect. 41.1 |  |  |  |
| --- | --- | --- | --- |
| Criteria | Compliance with Requirements? |  | Comments [ ] Manufacturer will not produce vacuum-cleaner or service cord. |
|  | Yes | No |  |
| Mechanical strength tester available |  |  | Manufacturer: Model: |
| Weight values: 150 lbf. (68.1 kg) |  |  |  |
| 170 lbf. (77.1 kg) |  |  |  |
| Calibrated within the last year |  |  | Last calibration date: Next calibration date: |
| Weight value permanently marked on weights |  |  |  |
| Clamps or other means available to securely hold cable sample under test conditions |  |  |  |
| Pulley or some other device for lifting specimen and weight evenly with no jerks. |  |  |  |

| Surface Marking Tester, Reference: UL 1581, Sect. 1690 |  |  |  |
| --- | --- | --- | --- |
| Criteria | Compliance with Requirements? |  | Comments |
|  | Yes | No |  |
| Surface marking tester available |  |  | Manufacturer: Model: |
| Calibrated within the last year |  |  | Last calibration date: Next calibration date: |
| Test block foot dimensions equal to 1 x 2 in. (25 x 50 mm) |  |  | Height of block: |
| Test block material is metal |  |  | [ ] Steel [ ] Brass [ ] Other: __________________ |
| Weight of test block equals 450 ± 5 gr (1 lbf ± 0.2 ozf) |  |  |  |
| Weight value permanently marked on test block |  |  |  |
| Felt available with thickness 0.047±0.005 in. (1.2±0.1 mm) |  |  |  |
| Felt securely attached to foot of test block |  |  |  |

| Ohmmeter, Reference: UL 1581, Sect. 220.1 |  |  |  |
| --- | --- | --- | --- |
| Criteria | Compliance with Requirements? |  | Comments |
|  | Yes | No |  |
| Ohmmeter available |  |  | Manufacturer: Model: |
| Calibrated within the last year |  |  | Last calibration date: Next calibration date: |
| Must be located in temperature controlled test room. |  |  |  |
| Minimum increment load scale no greater than 2% value read. |  |  | Record minimum increment of load scale: ohm for the Ranges: ohm respectively |

- N O T I C E  T O  M A N U F A C T U R E R -

The flame test equipment referenced on Pages 10 thru 13 are not required if you do not plan to conduct flame tests at the factory.  However, if you intend to conduct your own flame tests at the factory, we will require the information requested on Pages 10 thru 13.

Please check the appropriate space below:

We plan to conduct flame tests at the factory.  If so, please complete Pages 10 thru 13.

We do not plan to conduct flame tests at the factory.  If not, Pages 10 thru 13 do not need to be completed.

| Flame Chamber, Reference UL 1581, Sec. 1080 |  |  |  |
| --- | --- | --- | --- |
| Criteria | Compliance with Requirements? |  | Comments |
|  | Yes | No |  |
| Draft-free flame chamber available |  |  | Manufacturer: Model: |
| Flame chamber with interior volume of at least 140 ft3 or 4 m3 |  |  | Record dimensions: Width - Depth - Height - |
| Each linear interior dimension at least 24 in. or 610 mm |  |  |  |
| Windowed sash for access and viewing |  |  |  |
| Sash provided with an air-tight glove box for arm-and-hand access |  |  | [ ] Automatic burner positioning |
| Chamber interior visible without obstruction when sash is closed |  |  |  |
| Chamber fitted with an exhaust blower for pulling smoke and fumes |  |  |  |
| Tight sealing damper between chamber and exhaust blower |  |  |  |

| Flame Fuel and Gauges, Reference UL 1581, Sect. 1080 |  |  |  |
| --- | --- | --- | --- |
| Criteria | Compliance with Requirements? |  | Comments |
|  | Yes | No |  |
| Methane, propane, or natural gas available |  |  | Type of gas used: Grade: Purity: |
| Rate of calibration |  |  | ( ) Every 4 weeks, and ( ) Each new cylinder |
| Gauge to measure cylinder pressure |  |  |  |
| Gauge to measure pressure in supply line |  |  |  |
| Flowmeter present |  |  |  |
| Manometer/Pressure Gauge measuring to 800 mm of water, in 5 mm increments |  |  |  |

| Thermocouple Device For Flame Tester |  |  |  |
| --- | --- | --- | --- |
| Criteria | Compliance with Requirements? |  | Comments |
|  | Yes | No |  |
| Thermocouple device available |  |  | Manufacturer: Model: Thermocouple Type: |
| Calibrated within the last year |  |  | Last Calibration date: Next calibration date: |

| Flame Test Enclosure and Burner for Vertical Wire, VW-1, and Horizontal Flame Tests. Reference: UL 1581, Sect.1080, 1100 |  |  |  |
| --- | --- | --- | --- |
| Criteria | Compliance with Requirements? |  | Comments |
|  | Yes | No |  |
| Flame tester available |  |  | Manufacturer: Model: |
| Clamps provided no more than 17 in. apart to hold 18 in. wire specimen vertically in the enclosure |  |  |  |
| Clamps provided 8 in. apart to hold 10 in. wire specimen horizontally in the enclosure |  |  |  |
| Tirrell burner barrel length above air inlets 4 in. (102 mm) |  |  | Record actual length of barrel above air inlets: |
| Tirrell burner barrel inside diameter: 0.375 in. (9.5 mm) |  |  | Record actual inside diameter of barrel: |
| Air and gas adjustable to provide flame with an overall height of 4-5 in. (4-7/8 for VW-1) and a blue cone 1-1/2 in. (1-9/16 for VW-1) in height |  |  |  |
| Wedge dimensions in accordance with UL 1581, Figure 1080.2 |  |  | [ ] 20º angle provided |
| Wedge attached to metal plate with a hinge |  |  |  |
| Metal plate provided with rubber feet |  |  |  |
| Other required items available: Untreated cotton: 1/4 in. thick maximum Kraft paper to make indicator flag: Unreinforced 60 lb (94 g/m²), 5 mils (0.1 mm) thick gummed on one side, 1/2 in. wide Stop watch |  |  |  |

| Temperature Indicator For Flame Tester, ASTM D 5207 |  |  |  |
| --- | --- | --- | --- |
| Criteria | Compliance with Requirements? |  | Comments |
|  | Yes | No |  |
| Temperature indicator or potentiometer for Type K thermocouple capable of reading to 800ºC available |  |  | Manufacturer: Model: |
| Calibrated within the last year |  |  | Last calibration date: Next calibration date: |
| Thermal sensory (Type K thermocouple peened into top of copper slug) with slug dimensions 9.0 mm OD and 20.64 mm length |  |  | Record dimensions: Diameter - Length - |
| Ringstand with clamp for positioning thermal sensor available |  |  |  |
| Timer or stop watch provided to record temperature rise |  |  |  |

| Spark Tester, Reference: UL 1581, Sect. 900 |  |  |  |
| --- | --- | --- | --- |
| Criteria | Compliance with Requirements? |  | Comments |
|  | Yes | No |  |
| Spark tester available |  |  | Manufacturer: Model: |
| Calibrated within the last year |  |  | Last calibration date: Next calibration date: |
| Range of test voltage suitable for production |  |  | Range: 0 - kv |
| High voltage output connected to a single spark tester electrode. |  |  |  |
| Voltmeter required in circuit to indicate actual test potential at all times. |  |  |  |
| Equipment shall include a light, counter, or other system that gives a signal in the event of a fault. Fault indicator reset manually. |  |  |  |
| Visual or audible means of indicating test failure. |  |  |  |
| Spark tester suitable for maximum extrusion speed of production. |  |  | Maximum extrusion speed V = F x L x 0.55 Where: V = Velocity of wire in feet per minute: L = Minimum length of electrode in inches: F = Frequency of test voltage in Hertz: Hz Actual production speed: ft./min. |

| Spark Tester (continued), Reference: UL 1581, Sect. 900 |  |  |  |
| --- | --- | --- | --- |
| Criteria | Compliance with Requirements? |  | Comments |
|  | Yes | No |  |
| Transformer core and one end of secondary winding must be solidly connected to earth ground. |  |  |  |
| Electrode and enclosure covered by an earth grounded metal screen or equivalent guard to prevent shock hazard. |  |  |  |
| Conductors shall be earth grounded directly to the earth ground of the spark tester. Earth ground mentioned above connected to the service panel, grounding rod, or other acceptable means. |  |  | Type earth grounding provided: [ ] Grounded to service panel. [ ] Grounding rod. [ ] Other: __________________ |
| Wire used in earth grounding shall be a large AWG (14 AWG or larger), solid conductor that makes intimate contact with equipment at connection sites (e.g. paint breaking washers, etc.). Stranded conductor may be used if terminated with a connector. |  |  |  |
| Bottom of enclosure must be “U” or “V” shaped. Minimum width of enclosure approximately 1-1/2 inches greater than the diameter of the wire being tested. |  |  |  |
| Bead chains shall make intimate contact throughout the entire length of the wire being tested. |  |  | Bead chain configuration [ ] staggered [ ] unstaggered [ ] touching Bead Chain Details: bead diameter: mm chain length: mm longitudinal spacing: mm transverse spacing: mm vertical spacing: mm |
| The vertical spacing between beads in each chain shall not exceed the diameter of a bead. |  |  |  |
| Bead chain length must be appreciably greater than the depth of the enclosure. |  |  |  |
| Bead chains unbroken without missing chains. |  |  |  |

| Spark Tester - Brush Type, Reference: UL 1581, Sect. 900 |  |  |  |
| --- | --- | --- | --- |
| Criteria | Compliance with Requirements? |  | Comments |
|  | Yes | No |  |
| Spark tester available |  |  | Manufacturer: Model: |
| Calibrated within the last year |  |  | Last calibration date: Next calibration date: |
| Range of test voltage suitable for production |  |  | Range: 0 - kv |
| High voltage output connected to a single spark tester electrode. |  |  | Voltage frequency: Hz |
| Voltmeter required in circuit to indicate actual test potential at all times. |  |  |  |
| Equipment shall include a light, counter, or other system that gives a signal in the event of a fault. Fault indicator reset manually. |  |  |  |
| Visual or audible means of indicating test failure. |  |  |  |
| Spark tester suitable for maximum extrusion speed of production. |  |  | Maximum extrusion speed V = F x L x 0.55 Where: V = Velocity of wire in feet per minute: L = Minimum length of electrode in inches: F = Frequency of test voltage in Hertz: Hz Actual production speed: ft./min. |

| Spark Tester - Brush Type (continued), Reference: UL 1581, Sect. 900 |  |  |  |
| --- | --- | --- | --- |
| Criteria | Compliance with Requirements? |  | Comments: |
|  | Yes | No |  |
| Transformer core and one end of secondary winding must be solidly connected to earth ground. |  |  |  |
| Electrode and enclosure covered by an earth grounded metal screen or equivalent guard to prevent shock hazard. |  |  |  |
| Conductors shall be earth grounded directly to the earth ground of the spark tester. Earth ground mentioned above connected to the service panel, grounding rod, or other acceptable means. |  |  | Type earth grounding provided: [ ] Grounding to service panel. [ ] Grounding rod. [ ] Other: __________________ |
| Wire used in earth grounding shall be a large AWG (14 AWG or larger), solid conductor that makes intimate contact with equipment at connection sites (e.g. paint breaking washers, etc.). Stranded conductor may be used if terminated with a connector. |  |  |  |

| Buffing Machine, Reference: UL 1581, Sect. 420.6 |  |  |  |
| --- | --- | --- | --- |
| Criteria | Compliance with Requirements? |  | Comments |
|  | Yes | No |  |
| Buffing Machine Available |  |  | Manufacturer: Model: |
| Calibrated within the last year |  |  | Last calibration date: Next calibration date: |
| Specimen mounted in such a way that inside surface is splayed outward over a convex mount to provide even thickness after buffing |  |  |  |
| Abrasive Wheel: Approximately No. 36 grit |  |  | Grit size used: |
| Rotation per minute of wheel (2500 – 3500 r/min) |  |  | Record actual RPM: |
| Peripheral Speed of Wheel: 3000 – 5000 ft./min. (15 – 25 m/s) |  |  | Dia. of the wheel = Calculated Peripheral speed of wheel = RPM x  x Dia. of wheel = |
| Rate of feed: Slow, removing little compound at one cut, thereby not overheating specimen. |  |  |  |